This procedure is valid for the following chemistry analyzers:

| AU400/AU400e | AU640/AU640e |
| --- | --- |
| AU480 | AU680 |
| AU600 | AU2700/AU5400 |

| Prepared By | Date Adopted | Supersedes Procedure # |
| --- | --- | --- |

| Review Date | Revision Date | Signature |
| --- | --- | --- |

| Distributed to | # of Copies | Distributed to | # of Copies |
| --- | --- | --- | --- |

# PRINCIPLE:

Lithium is widely used in the treatment of manic depressive psychosis.  Administered as Lithium Carbonate, it is completely absorbed by the gastro-instestinal tract, peak serum levels occur 2 to 4 hours after an oral dose.  The half life in serum is 48 to 72 hours, and it is cleared through the kidneys (excretion parallels that of sodium).  Reduced renal function can prolong clearance time.1,2

Lithium acts by enhancing the uptake of neurotransmitters, which produces a sedative effect on the central nervous system.  Serum lithium concentrations are carried out essentially to ensure compliance and to avoid toxicity.1,2

Early symptoms of intoxication include apathy, sluggishness, drowsiness, lethargy, speech difficulties, irregular tremors, myoclonic twitchings, muscle weakness, and ataxia.  Levels higher than 1.5 mmol/L (12 hours after a dose) indicate a significant risk of intoxication.1,2

# INTENDED USE:

System reagent for the quantitative determination of lithium concentrations in human serum, or plasma on Beckman Coulter AU Clinical Chemistry analyzers.

# METHODOLOGY

The Beckman Coulter Lithium Assay is a spectrophotometric method, which can be readily adapted to automated clinical chemistry analyzers.  Lithium present in the sample reacts with a substituted porphyrin compound at an alkaline pH.  The resulting change in absorbance is directly proportional to the concentration of lithium in the sample.1

Lithium can be determined by several methods that include atomic absorption spectrophotometry, flame emission photometry, or ion-selective electrodes.  These methods require specific and often dedicated instrumentation.1

# SPECIMEN:

## Patient / Sample Preparation:

No special preparation for the patient is required. The patient’s clinical condition and dosage regimen may influence the sample collection time.

## Sample Collection Time:

It is recommended that a standardized 12hr-post dose serum lithium concentration be used to assess adequate therapy.  Peak concentration is reached 2 to 4 hours after oral dosing.1,2,3

Additional instructions for patient preparation as designated by this laboratory:

## Type:

Serum or EDTA plasma is the recommended specimen.  Whole blood cannot be used.

Additional type conditions as designated by this laboratory:

## Handling Conditions:

Serum or plasma should be separated from cells if storage of more than 4 hours is anticipated.1,2,3

Lithium samples are stable for one week when stored at 2 - 25(C and stable indefinitely when frozen < -20(C.8

Human serum or plasma samples should be handled and disposed of as if they were potentially infectious.

Additional handling conditions as designated by this laboratory:

# EQUIPMENT AND MATERIALS:

## Equipment:

Beckman Coulter AU400/AU400e, AU480, AU600, AU640/AU640e, AU680, AU2700, and AU5400 analyzers.

## Materials:

Beckman Coulter System Lithium

Final concentration of active ingredients:

Sodium hydroxide				0.5 mol/L

EDTA							50 (mol/L

Substituted Porphyrin			15 (mol/L

Reagent storage location in this laboratory:

Test tubes 12 -16 mm in diameter or sample cups 	(Cat No.AU1063).

Storage location of test tubes or sample cups in this laboratory:

Lithium Standard 				(provided with kit OSRT056)

The Beckman Coulter Lithium standard contains 1.0 mmol/L of lithium chloride with preservative.  The standard is prepared from a stock lithium solution (lithium dissolved in HCl) diluted in physiological saline (0.5 g/L).9

Storage location of the calibrator in this laboratory:

### Preparation

The Beckman Coulter Lithium reagent is ready to use.  No preparation is needed.

The Beckman Coulter Lithium Standard is provided as a liquid and is ready to use.

### Precautions:

1.	The Beckman Coulter Lithium Assay and Standard are for in vitro diagnostic use.

2.	Do not ingest.  Harmful if swallowed.  Avoid contact with skin and eyes.  If spilled, thoroughly wash affected areas with water.

3.	Assay and standard contain sodium azide (0.1%), which may react with lead and copper plumbing to form highly explosive metal azides.  If waste is discarded down the drain, flush it with a large volume of water to prevent azide buildup.  Dispose of properly in accordance with local regulations.

4.	Do not use the reagent kit or standard after the expiration date.

### Storage and Stability:

1.  The unopened reagent and standard are stable until the expiration date printed on the label when stored at 2 - 8°C.

- Opened reagent is stable for 14 days when stored in the refrigerated compartment of the analyzer.  If the opened reagent is capped when not in use and stored at 2 - 8°C, the reagent is stable until the expiration date printed on the label.

- Opened standard is stable until the expiration date printed on the label when stored at 2 - 8°C, provided the bottle is stored capped.

Improper storage of reagents or calibrators can affect assay performance. Stability depends on handling reagents or calibrators as directed.

Additional storage requirements as designated by this laboratory:

### Indications of Deterioration:

Color of reagent is light purple.  Discoloration of the reagent or calibrators, visible signs of microbial growth, turbidity, or precipitation in reagent or calibrators may indicate degradation and warrant discontinuance of use.

# PERFORMANCE PARAMETERS:

The following performance characteristics represent total system performance and should not be interpreted to refer only to reagents. Studies were performed on the Beckman Coulter AU analyzer series.  Results may vary due to analyzer-to-analyzer differences.

## Precision

Estimates of precision, based on CLSI EP5-A recommendations7, are consistent with typical performance. The within run precision is less than 3% and total precision is less than 5%.  Assays of control materials were performed and this data produced following CLSI guidelines.6

| N=80 | Within run | Total |
| --- | --- | --- |
| Mean, mmol/L | SD | CV% | SD | CV% |
| 0.57 | 0.005 | 0.9 | 0.011 | 1.9 |
| 1.83 | 0.012 | 0.7 | 0.024 | 1.3 |

## Comparison

A comparison of this Beckman Coulter lithium method (Method 1) vs NOVA ISE (Method 2) was run on an AU2700/AU5400 analyzers utilizing 55 patient serum samples.  The following data demonstrates representative performance on AU analyzers.6

| Slope | 1.01 |
| --- | --- |
| Intercept (mmol/L) | 0.010 |
| Correlation Coefficient | 0.9959 |
| Range of patient results | 0.11 – 1.72 mmol/L |
| Number of Samples | 55 |

# CALIBRATION:

Perform a one-point calibration (AB) using a water blank (blue rack) and the Beckman Coulter lithium standard provided in the kit.  Calibration parameters are set to prepare the calibration curve.  Refer to analyzer User’s Guide or Method Parameter sheets for analyzer settings.

## Calibration Stability

The calibration is stable for 7 days.  Recalibrate as indicated by control results or whenever a new lot of reagent is used.  Calibration stability may vary from laboratory to laboratory depending on the following: handling of reagents, maintenance of analyzer, adherence to operating procedures, establishment of control limits, and verification of calibration.

# QUALITY CONTROL:

During operation of the Beckman Coulter AU analyzer at least two levels of control material should be tested routinely with each group of unknown samples or a minimum of once a day.  Controls should be performed after calibration, with each new lot of reagent, and after specific maintenance or troubleshooting steps described in the appropriate User’s Guide.  Quality control testing should be performed in accordance with regulatory requirements and individual laboratory’s standard procedures.  If more frequent verification of test results is required by the operating procedures within a laboratory, those requirements should be met.

# PARAMETERS:

The Beckman Coulter Lithium parameters include a Pre-Dilution Rate of 10 that should be noted on the parameter sheets.

A complete list of test parameters and operating procedures can be found in the appropriate User’s Guide and at www.beckmancoulter.com.

# CALCULATIONS:

Results are calculated automatically by the analyzer.  No additional manipulation of data is required.

# REPORTING RESULTS:

## Reference Ranges:1,2

12 hour post dose trough concentration:		1.0 – 1.2 mmol/L

Values > 1.5 mmol/L 12 hours post dose indicates a significant risk of intoxication.

The quoted values should serve as a guide only.  It is recommended that each laboratory verify this range or derive a reference interval for the population it serves.5

Expected reference ranges in this laboratory:

## Procedures for Abnormal Results

The laboratory must define procedures to be used in reporting high concentration (toxic) results to the patient’s physician.

Abnormal results are flagged by the listed analyzers according to the normal values entered by the user into the analyzer parameters.

## Reporting Format:

Results are automatically printed out for each sample in mmol/L at 37°C.

Additional reporting information as designated by this laboratory:

# LIMITATIONS:

The Beckman Coulter Lithium Assay is linear from 0.10 – 5.0 mmol/L.

Alteration of reagents, use of analyzers without appropriate capabilities, or other failure to follow instructions as set forth in this protocol or the package insert can affect performance characteristics and stated or implied labeling claims.

## Interfering Substances

The reagent is light sensitive and will absorb atmospheric carbon dioxide.  It is recommended that the reagent be stored capped and in a dark container when not in use for prolonged periods of time (eg. overnight).

Studies to determine the level of interference from other cations normally present in serum were carried out in the presence of a lithium concentration of approximately 1 mmol/L and the following results were obtained:

No significant interference (< 5% deviation from assigned lithium concentration) from the following were observed with this method:

| Sodium | Up to 200 mmol/L |
| --- | --- |
| Potassium | Up to 8.00 mmol/L |
| Calcium | Up to 4.00 mmol/L (16 mg/dL) |
| Magnesium | Up to 2.00 mmol/L (4.86 mg/dL) |
| Iron | Up to 200 µmol/L (1117 µg/dL) |
| Zinc | Up to 250 µmol/L (1625 µg/dL) |
| Copper | Up to 250 µmol/L (1588 µg/dL) |

## Sensitivity

The lowest level of detection for this method was determined by assaying 11 replicates of a serum substitute (SERASUB™) that does not contain lithium.  When run as recommended, the lowest detection limit is 0.04 mmol/L.

# REFERENCES:

1.	Tietz. Fundamentals of Clinical Chemistry, Sixth Edition. Saunders Elservier Inc., 2008 pg 555,556,868.

2.	Amdisen A. “Serum Lithium determinations for Clinical Use.” Scand Jnl Clin Lab Invest. 1967; 20:104-8.

3.	Young DS. “Effects of Pre-analytical Variables on Clinical Laboratory Tests.” Second Edition. 1997, pg 3-360.

4.	Tietz NW. “Blood Gases and Electrolytes in Fundamentals of Clinical Chemistry, Philadelphia W.B. Saunders Co., 1976 pg 899-901.

5.	Wachtel M et al. “Creation and Verification of Reference Intervals.” Laboratory Medicine. 1995; 26:593-7.

6.	Data is on file for specific AU analyzers.

7.	Clinical and Laboratory Standards Institute. User evaluation of Precision Performance of Clinical Laboratory Devices.  CLSI:2004, CLSI Publication EP5-A2.

8.	CAP Patient Prep & Specimen Handling Fascicle IV: Therapeutic Drug Monitoring/Toxicology. References:  Tietz NW. “Blood gases and electrolytes.” In Fundamentals of Clinical Chemistry. Philadelphia. W.B. Saunders Co., 1976, pp 899-901.

9.	ThermoTrace Lithium Standard package insert. ThermoTrace, Noble Park, Victoria, Australia. 1-800-558-9115.  P/N:P1669040.03 revised September 2001.