This procedure is valid for the following chemistry analyzers:

| AU400/AU400e | AU640/AU640e |
| --- | --- |
| AU480 | AU680 |
| AU600 | AU2700/AU5400 |

| Prepared By | Date Adopted | Supersedes Procedure # |
| --- | --- | --- |

| Review Date | Revision Date | Signature |
| --- | --- | --- |

| Distributed to | # of Copies | Distributed to | # of Copies |
| --- | --- | --- | --- |

# PRINCIPLE:

Cocaine is a central nervous system stimulant that is extracted from the coca plant.  As a drug of abuse, it is self-administered is a variety of ways, including inhalation and intravenous injections.  Cocaine base can be smoked in a form that is commonly known as "crack".  Cocaine is rapidly absorbed, especially when smoked.  While all forms are potentially addicting, "crack" is especially likely to lead to dependence because of its more rapid and heightened effect on the abuser.2

Excretion rate patterns vary with the mode of administration and from individual to individual.  Cocaine is almost completely metabolized, primarily in the liver, with only about one percent excreted in the urine unchanged. Most cocaine is eliminated as benzoylecgonine, the major metabolite of cocaine.  Cocaine is also excreted in relatively lesser amounts as ecgonine methylester, and ecgonine.  Cocaine metabolites may be detected in urine for up to a couple of days after cocaine is used.  Benzoylecgonine can be detected in urine within four hours after cocaine inhalation and remain detectable in concentrations greater than 1000 ng/mL for as long as 48 hours.3-6

The Emit( II Plus Cocaine Metabolite Assay tests for benzoylecgonine, the major metabolite of cocaine, in human urine.  Positive results for samples containing other compounds structurally unrelated to benzoylecgonine have not been observed.

Methods historically used for detecting benzoylecgonine in biological fluids include high-performance liquid chromatography, gas-liquid chromatography, and enzyme immunoassay.7-9

# INTENDED USE:

The Emit( II Plus Cocaine Metabolite Assay is a homogeneous enzyme assay with a 150 ng/mL or 300 ng/mL cutoff (SAMSHA initial test cutoff level).1  This assay is intended for use in the qualitative and semi-quantitative analysis of benzoylecgonine (cocaine metabolite) in human urine,  Emit( II Plus assays are designed for use with multiple Beckman Coulter AU analyzers.

# METHODOLOGY

The Emit® II Plus Cocaine Metabolite Assay is a homogeneous enzyme immunoassay technique used for the analysis of specific compounds in human urine.10 The assay is based on competition between drug in the sample and drug labeled with the enzyme glucose-6-phosphate dehydrogenase (G6PDH) for antibody binding sites.  Enzyme activity decreases upon binding to the antibody, so the drug concentration in the sample can be measured in terms of enzyme activity. Active enzyme converts nicotinamide adenine dinucleotide (NAD) to NADH, resulting in an absorbance change that is measured spectrophotometrically.  Endogenous serum G6PDH does not interfere because the coenzyme NAD functions only with the bacterial (Leuconostoc mesenteroides) enzyme employed in the assay.

The Emit( II Plus Cocaine Metabolite Assay provides only a preliminary analytical test result.  A more specific alternative chemical method must be used to obtain a confirmed analytical result.  Gas chromatography/mass spectrometry (GC/MS) is the preferred confirmatory method2 but other chemical confirmation methods are available.  Clinical consideration and professional judgment should be applied to any drug of abuse test result, particularly when preliminary positive results are used.

# SPECIMEN:

## Patient / Sample Preparation:

None required.

Additional instructions for preparation as designated by this laboratory:

## Type:

Urine samples are the recommended specimen type.

Additional type conditions as designated by this laboratory:

## Handling Conditions:

Urine specimens may be collected in plastic (i.e., polypropylene, polycarbonate, polyethylene) or glass containers. Some plastics, other than those listed, can adsorb certain drugs.

If not analyzed immediately, specimens may be stored at room temperature (15-25C) for up to 7 days following collection.  After 7 days, specimens should be stored frozen (< -20C).  Frozen specimens must be completely thawed and mixed thoroughly prior to analysis.

Specimens with high turbidity should be centrifuged before analysis.

The recommended pH range for urine specimens is 3.0-11.0. However, benzoylecgonine can degrade upon prolonged exposure to pH levels greater than 9.0.

Adulteration of the urine specimen may cause erroneous results. If adulteration is suspected, obtain another specimen.

Human urine specimens should be handled and treated as if they are potentially infectious.

Additional handling conditions as designated by this laboratory:

# EQUIPMENT AND MATERIALS:

## Equipment:

Beckman Coulter AU400/AU400e, AU480, AU600, AU640/AU640e, AU680, AU2700, and AU5400 analyzers.

## Materials:

Emit( II Plus Cocaine Metabolite Assay

Antibody/Substrate Reagent 1 -- Sheep polyclonal antibodies reactive to benzoylecgonine, glucose-6-phosphate, nicotinamide adenine dinucleotide, bovine serum albumin, preservatives, and stabilizers

Enzyme Reagent 2 -- Benzoylecgonine labeled with glucose-6-phosphate dehydrogenase, HEPES buffer, bovine serum albumin, preservatives and stabilizers

Reagent storage location in this laboratory:

Test tubes 12 -16 mm in diameter or sample cups 	(Cat No. AU1063).

Storage location of test tubes or sample cups in this laboratory:

Emit( Calibrator/Control products are packaged individually and sold separately.

Emit( Calibrator/Control Level 0                          	Cat No.  9A509

Emit( Calibrator/Control Level 2 (150 ng/mL)     	Cat No.  9A549

Emit( Calibrator/Control Level 3 (300 ng/mL)     	Cat No.  9A569

Emit( Calibrator/Control Level 4 (500 ng/mL)      	Cat No.  9A589

Emit( Calibrator/Control Level 5 (1000 ng/mL)     	Cat No.  9A609

Note:  The Emit( Calibrator/Controls contain stated concentrations of benzoylecgonine (ng/mL) for calibration of this assay.  Refer to package insert for concentration listings.

Storage location of the calibrators in this laboratory:

### Preparation

The Emit® II Plus Cocaine Metabolite Assay reagents are packaged in a ready to use liquid form and may be used directly from the refrigerator.

Note: Reagents 1 and 2 are sold as a matched set. They should not be interchanged with components of kits with different lot numbers.

The Emit® Calibrators/Controls are packaged in a ready to use liquid form and may be used directly from the refrigerator.  Close the calibrator bottles when not in use. Caps must always be replaced on the original containers.

### Precautions:

1.	The Emit® II Plus Cocaine Metabolite Assay and Calibrator / Controls are for in vitro diagnostic use.

- Reagent 1 contains nonsterile sheep antibodies.  Reagent 2 contains non-sterile mouse antibodies.  Non-sterile bovine serum albumin is found in both Reagent 1 and 2.

3.	No known test method can offer complete assurance that products derived from human sources or inactivated microorganisms will not transmit infection.  Reagents, calibrators, and human specimens should be handled using prevailing good laboratory practices to avoid skin contact or ingestion.

4.	Do not use the reagents or calibrators after the expiration date.

5.	This Emit( II Plus Cocaine Metabolite Assay is qualified for use only with the Emit( Calibrators listed in the Calibrator section.

### Storage Requirements:

Any reagents not loaded in the reagent refrigerator on the analyzer or any calibrators not in use should be stored at 2-8°C, upright, and with caps tightly closed.  Do not freeze reagents or calibrators.  Avoid exposure to temperatures above 32°C for prolonged periods of time.

Unopened reagents and calibrators are stable until the expiration date printed on the label if stored as directed.  Refer to Assay Methodology Sheets for additional on-board stability information.

Improper storage of reagents or calibrators can affect assay performance.  Stability depends on handling reagents and calibrators as directed.

Additional storage requirements as designated by this laboratory:

### Indications of Deterioration:

Discoloration (especially yellowing) of the reagents or calibrators, visible signs of microbial growth, turbidity, or precipitation in reagent or calibrator may indicate degradation and warrant discontinuance of use.

# PERFORMANCE PARAMETERS:

The following performance characteristics represent total system performance and should not be interpreted to refer only to reagents. Studies were performed on the Beckman Coulter AU analyzer series.  Results may vary due to analyzer-to-analyzer differences.  Positive results were confirmed by GC/MS.

## Precision

Within run precision was performed and calculated according to Clinical and Laboratory Standards Institute (CLSI EP5-A) by running 2 replicates of the cutoff Calibrator/Control and positive and negative controls twice a day for 20 days (N=80).  Total precision was calculated from this data.  Results for these studies are summarized in the following tables for respective cutoffs.

Cocaine Metabolite (150 ng/mL Cutoff)

|  | Within Run Precision | Total Precision |
| --- | --- | --- |
|  | Cutoff Cal | Control 75% | Control 125% | Cutoff Cal | Control 75% | Control 125% |
| Mean mAU/min | 314 | 293 | 329 | 314 | 293 | 329 |
| SD | 2.1 | 2.2 | 2.8 | 3.6 | 3.4 | 4.3 |
| %CV | 0.6 | 0.6 | 0.7 | 1.0 | 1.0 | 1.1 |

Cocaine Metabolite (300 ng/mL Cutoff)

|  | Within Run Precision | Total Precision |
| --- | --- | --- |
|  | Cutoff Cal | Control 75% | Control 125% | Cutoff Cal | Control 75% | Control 125% |
| Mean mAU/min | 393 | 283 | 402 | 393 | 283 | 402 |
| SD | 3.5 | 1.1 | 3.4 | 5.3 | 5.6 | 5.1 |
| %CV | 0.9 | 1.1 | 0.9 | 1.4 | 1.5 | 1.3 |

## Comparison

Clinical urine specimens were tested using the Emit( II Plus Cocaine Metabolite Assay at the 150 ng/mL cutoff and 300 ng/mL cutoff on an Beckman Coulter AU analyzer and the corresponding Emit( II assay on the SYVA(-30R Biochemical System.  Specimens positive by either method were confirmed by GC/MS analysis.  The results are summarized below showing the number of positive and negative results identified and the percent agreement between analyzers.

| Assay | Positive | Negative | %Agreement |
| --- | --- | --- | --- |
| Cocaine Metabolite 150 | 62 | 57 | 95.2 |
| Cocaine Metabolite 300 | 37 | 85 | 97.6 |

## Analytical Recovery

Negative human urine specimens were spiked with concentrations of benzoylecgonine.  Specimens spiked with drug concentrations lower than the cutoff concentration were analyzed qualitatively and correctly identified as negative 100% of the time.  Specimens spiked with drug concentrations greater than the cutoff were correctly identified as positive 100% of the time.  Results from semi-quantitative analysis of the specimens are listed below:

| Concentration (ng/mL) | Mean (ng/mL) |
| --- | --- |
| 100 | 94 |
| 225 | 244 |
| 270 | 288 |
| 330 | 311 |
| 450 | 470 |
| 750 | 833 |

# CALIBRATION:

## Qualitative Analysis

Perform a one-point calibration (AB) using a water blank (blue rack) and the appropriate EMIT Calibrator / Control for the desired cutoff: Level 2 = 150 ng/ml or Level 3 = 300 ng/ml.  Refer to Analyzer Specific Protocol for calibration set-point options and analyzer settings.

Three options are available for Qualitative Calibration:

Option 1:  On the “Specific Test Parameters” menu, the “General Tab”, program the Correlation B factor as 0.0.  On the same screen, under the “Range Tab” program the Value/Flag Level H as 999999.  Under Calibration Specific Parameters menu set the Calibration type to MB.  Blank the test using the blue rack.  The cutoff calibrator (150 or 300) is run in a white rack.  Each sample response is compared to the cutoff calibrator response to determine if the sample is positive or negative.  Positive samples will not be flagged.  Comparison of sample responses is a manual process.

Option 2:  On the “Specific Test Parameters” menu, the “General Tab”, program the Correlation B factor as 0.0.  On the same screen, under the “Range Tab” program the Value/Flag Level H as 100.  Under Calibration Specific Parameters menu set the Calibration type to AB with Formula as Y=ax+b.  The Conc. for the calibrator should be entered as 100.  Blank the test using the blue rack.  Calibrate by placing the designated cutoff calibrator (150 or 300) in the assigned position in the calibration rack (yellow rack). Positive samples will be flagged (P) and will printout as greater than or equal to 100.

Option 3:  On the “Specific Test Parameters” menu, the “General Tab”, program the Correlation B factor as -100.  On the same screen, under the “Range Tab” program the Value/Flag Level H as 0.0.  Under Calibration Specific Parameters menu set the Calibration type to AB with Formula as Y=ax+b.  The Conc. for the calibrator should be entered as 100.  Blank the test using the blue rack.  Calibrate by placing the designated cutoff calibrator (150 or 300) in the assigned position in the calibration rack (yellow rack). Positive samples will be flagged (P) and will printout as greater than or equal to zero.

## Semi-Quantitative Analysis

Perform a multi-point calibration (4AB) using a water blank (blue rack) and the EMIT Calibrator / Controls:  Level 2, Level 3, Level 4, and Level 5.  Calibration parameters are set to prepare the calibration curve.  Refer to analyzer User’s Guide or Analyzer Specific Protocol sheets for analyzer settings.

## Calibration Stability

Studies have shown the median calibration stability to be at least 14 days.  Recalibrate as indicated by control results or with a new lot of reagent.  Calibration stability may vary from laboratory to laboratory depending on the following: handling of reagents, maintenance of analyzer, adherence to operating procedures, establishment of control limits, and verification of calibration.

Note:  When using a new set of reagents with the same lot number, recalibration may not be required.  Validate the system by assaying controls.

# QUALITY CONTROL:

During operation of the Beckman Coulter AU analyzer at least two levels of control material should be tested a minimum of once a day.  Controls should be performed after calibration, with each new lot of reagent, and after specific maintenance or troubleshooting steps described in the appropriate User’s Guide.  Quality control testing should be performed in accordance with regulatory requirements and individual laboratory’s standard procedures. If more frequent verification of test results is required by the operating procedures within your laboratory, those requirements should be met.

## Qualitative Analysis

Validate the calibration by assaying controls.  Ensure that the result from the negative control is negative (or lower) relative to the Calibrator/ Control set point.  Ensure that the result from the positive is positive (or higher) relative to the Calibrator/Control set point.  Once the calibration is validated, run urine specimens.

## Semi-Quantitative Analysis

Validate the calibration by assaying controls.  Ensure that control results fall within acceptable limits as defined by the testing facility.  Once the calibration is validated, run urine specimens.

# PARAMETERS:

A complete list of test parameters and operating procedures can be found in the appropriate User’s Guide and at www.beckmancoulter.com.  The Analyzer Specific Protocol Sheets may also be used.

# CALCULATIONS:

None required.

# REPORTING RESULTS

## Reference Ranges:

No reference ranges are defined for drugs of abuse testing.

Expected reference ranges in this laboratory:

## Procedures for Abnormal Results:

The laboratory must define procedures to be used in reporting high concentration (toxic) results to the patient’s physician.

Abnormal results are flagged by the listed analyzers according to the normal values entered by the user into the instrument parameters.

## Reporting Format:

Semiquantitative results are automatically in ng/mL at 37C.

### Interpretation of Results

Qualitative Analysis -- When the Emit( II Plus Cocaine Metabolite Assay is used as a qualitative assay, the amount of drug and metabolites detected by the assay in any given sample cannot be estimated.  The assay results distinguish positive from negative samples only.  The Emit® Calibrator/Control Cutoff as designated by the testing facility, which contains either a concentration of 150 ng/mL or of 300 ng/mL, is used as a reference for distinguishing “positive” from “negative” specimens.

### Positive Results:  A specimen that gives a result equal to or higher than the Calibrator/Control set point is interpreted as positive: The specimen contains benzoylecgonine.

### Negative Results:  A specimen that gives a result lower than the Calibrator/Control set point is interpreted as negative: Either the specimen does not contain benzoylecgonine or benzoylecgonine is present in concentrations below the cutoff level for this assay.

Semi-Quantitative Analysis -- When used semi-quantitatively, the Emit( II Plus Cocaine Metabolite Assay, yields approximate, cumulative concentrations of the drug and metabolites detected by the assay. The semi-quantitation of positive results enables laboratories to determine an appropriate dilution of the specimen for confirmation by GC/MS.  Semi-quantitation also permits laboratories to establish quality control procedures and assess control performance.

Immunoassays that produce a single result in the presence of multiple detectable components cannot fully quantitate the concentration of individual components.  Interpretation of results must take into account that urine concentrations can vary extensively with fluid intake and other biological variables.  A more specific alternative chemical method must be used to obtain a confirmed analytical result.

Additional reporting information as designated by this laboratory:

# LIMITATIONS:

- This Emit® II Plus Cocaine Metabolite Assay is designed for use only with human urine.

- A positive result from this assay indicates the presence of cocaine metabolites but does not indicate or measure intoxication.

- Boric acid is not recommended as a preservative for urine.

- Other substances and/or factors not listed (e.g. technical or procedural errors) may interfere with the test and cause false results.

- Interpretation of results must take into account that urine concentrations of cocaine metabolites can vary extensively with fluid intake and other biological variables.

- Immunoassays that produce a single result in the presence of a drug and its metabolites cannot fully quantitate the concentration of individual components.

## Sensitivity

The minimum detection limit of the Emit® II Plus Cocaine Metabolites assay is less than 20 ng/mL.  This level represents the lowest concentration of cocaine that can be distinguished from 0 ng/mL with a confidence level of 95%.

## Specificity

The Emit® II Plus Cocaine Metabolite Assay detects benzoylecgonine, the major metabolite of cocaine, in human urine.

The following table lists the concentrations of compounds that produce a result that is approximately equivalent to the Emit® Calibrator/Control Level 2; 150 ng/mL and Level 3; 300 ng/mL calibrator/control cutoffs, respectively.  Each concentration represents the reactivity level for the stated compound when it is added to a negative urine specimen.  These concentrations are within the range of levels found in urine following use of the compound or, in the case of metabolites, the parent compound.  If a specimen contains more than one compound detected by the assay at lower concentrations than those listed in table, these may combine to produce a rate approximately equivalent to or greater than that of the cutoff calibrator.

| Compound | Concentration (ng/mL) at 150 ng/mL Cutoff | Concentration (ng/mL) at 300 ng/mL Cutoff |
| --- | --- | --- |
| Cocaine | 18-53 | 40-119 |
| Ecgonine | 2-6 | 7-20 |

The following table gives a list of the compounds that produce a negative result by the Emit® II Plus Cocaine Metabolite Assay.  Specificity testing was performed at the 150 ng/mL cutoff, which represents the greatest potential for cross-reactivity.  Positive results for compounds structurally unrelated to cocaine metabolite have not been observed.

| Compound | Concentration Tested (ng/mL) at the 150 ng/mL Cutoff |
| --- | --- |
| Acetaminophen | 1000 |
| (-Acetyl-N,N-dinormethadol (dinor LAAM) | 25 |
| L-(-Acetylmethadol (LAAM) | 25 |
| N-Acetylprocainamide (NAPA) | 400 |
| Acetylsalicylic Acid | 1000 |
| Amitriptyline | 1000 |
| Buprenorphine | 1000 |
| Caffeine | 1000 |
| Compound | Concentration Tested (ng/mL) at the 150 ng/mL Cutoff |
| Cimetidine | 1000 |
| Clomipramine | 2.5 |
| Clonidine | 1000 |
| Codeine | 500 |
| Cotinine | 100 |
| Cyclobenzaprine | 1000 |
| Desipramine | 800 |
| Dextromethorphan | 1000 |
| Diphenhydramine | 1000 |
| Doxepin | 1000 |
| 2-Ethylidene-1,5-dimethyl-3,3-diphenylpyrrolidine (EDDP) | 1000 |
| Fluoxetine | 1000 |
| Glutethimide | 500 |
| Ibuprofen | 1000 |
| Ketamine | 100 |
| Ketorolac Tromethamine | 1000 |
| Lormetazepam | 1 |
| LSD | 0.01 |
| Meperidine | 1000 |
| Methadone | 1000 |
| Methaqualone | 1500 |
| Compound | Concentration Tested (ng/mL) at the 150 ng/mL Cutoff |
| Morphine | 1000 |
| Naproxen | 1000 |
| Nortriptyline | 1000 |
| Oxazepam | 300 |
| Phencyclidine | 1000 |
| Phenytoin | 1000 |
| Promethazine | 1000 |
| Propoxyphene | 1000 |
| Ranitidine | 1000 |
| Scopolamine | 500 |
| Secobarbital | 1000 |
| 11-nor-(9-THC-9-COOH | 100 |
| Thioridazine | 100 |
| Tramadol | 1000 |
| Tyramine | 100 |
| Zidovudine (AZT) | 2000 |
| Zolpidem | 100 |

# REFERENCES:

1.	Mandatory Guidelines for Federal Workplace Drug Testing Programs: Final Guidelines: Effective September 1, 1994.

2.	Hawks RL, Chiang CN, eds. Urine Testing for Drugs of Abuse. Rockville, MD: National Institute on Drug Abuse (NIDA), NIDA research monograph 73. Department of Health and Human Services; 1986.

3.	Wyngaarden JB, Smith LH Jr, eds. Cecil Textbook of Medicine. Philadelphia, PA: WB Saunders Co; 1988:54–55.

4.	Pearman K. Cocaine: A review. J Laryngol Otol. 1979; 93:1191–1199.

5.	Stewart DJ. et al. Cocaine metabolism: Cocaine and norcocaine hydrolysis by liver and serum esterases. Clin Pharmacol Ther. 1978; 25:464–468.

6.	Hamilton HE et al. Cocaine and benzoylecgonine excretion in humans. J Forensic Sci Sas. 1977; 22:697–707.

7.	Van Dyke C et al. Urinary excretion of immunologically reactive metabolite(s) after intranasal administration of cocaine as followed by enzyme immunoassay. Clin Chem. 1977; 23:241–244.

8.	Ellenhorn MJ, Barceloux DG. Medical Toxicology. New York, NY: Elsevier Science Publishing Company, Inc. 1988:644–659.

9.	Kogan MJ et al. Quantitative determination of benzoylecgonine and cocaine in human biofluids by gas-liquid chromatography. Anal Chem. 1977; 49:1965–1969.

10.	Lindgren JE. Guide to the analysis of cocaine and its metabolites in biological material. J Ethnopharmacol. 1981; 3:337–351.

11. Oellerich, M.  Enzyme immunoassays in clinical chemistry: present status and trends. J Clin Chem Clin Biochem. 1980; 18:197-208.